Supplementary Information

Hybrid Bayesian Inference and Reinforcement Learning for Autonomous pH Adjustment in Diverse Chemical Systems

Siyuan Zhang1, Lin Huang1†, Tianci Shi1†, Yibo Xie1,4, Jing Yang2*, Yibin Jiang1,2,3*, Cheng Wang1,2,3,4*

1iChem, State Key Laboratory of Physical Chemistry of Solid Surfaces, College of Chemistry and Chemical Engineering, Xiamen University, Xiamen 361005, P. R. China.

2College of Chemistry and Chemical Engineering, Xiamen University, Xiamen 361005, P. R. China.

3Innovation Laboratory for Sciences and Technologies of Energy Materials of Fujian Province (IKKEM), Xiamen 361005, P. R. China.

4Institute of Artificial Intelligence, Xiamen University, Xiamen 361005, P. R. China.

†These authors contributed equally. *Corresponding author email:

j.yang@xmu.edu.cn, yibinjiang@xmu.edu.cn, wangchengxmu@xmu.edu.cn.

Table of Contents

Section 1 Introduction to the Automatic pH-Adjustment Device	2

Section 2 Chemical Reagents	3

Section 3 Experiment Validation	4

Experiment 1 Benchmarking Three Control Policies in a Robotic pH Adjustment Task	4

Experiment 2 Wastewater Neutralization	7

Experiment 3 Isolation of Casein by Adjusting pH to the Isoelectric Point	8

Experiment 4 Determination of the Stability Constant for the SSA–Cu(II) Complex	10

Section 4 Bayesian Updating and Reinforcement Learning	13

Bayesian Parameter Updating	13

1. Sampling	13

2. Equilibrium prediction	14

3. Likelihood calculation	15

4. Resampling	15

5. Posterior updating	16

Determining pH-Adjustment Volume Using Bayesian Parameter Updates and Control Logic	16

1. Current pH error magnitude (ΔpHerr)	17

2. Observed pH response to the previous addition (r)	18

3. Posterior uncertainty (u)	18

4. Effective buffering level (B)	19

5. Theoretical required volume (Vreq)	19

Imitation Learning	21

Optimizing Neural Network Parameters via Reinforcement Learning	23

Parameter update process	23

Reward Function Design	25

1. Dense reward	25

2. Step penalty	26

3. Overshoot penalty	26

4. Incorrect reagent penalty	26

5. Overshoot-related volume penalty and bonus	27

6. Terminal reward	27

Training Procedure	28

ΔpHerr

r

u

B

Vreq

ΔpHerr

r

u

B

Vreq

ΔpHerr

r

u

B

Vreq

(rtdense)

Section 1 Introduction to the Automatic pH-Adjustment Device

As shown in Figure S1, the solution to be adjusted was contained in a three-necked flask. The left neck was connected to a nitrogen gas inlet, which generated bubbles in the solution to facilitate stirring. The central neck was fitted with a pH probe to monitor the solution’s pH value. The right neck was connected to the reagent delivery tubing.

During the pH-adjustment, the control algorithm determined the reagent volume/type, and a syringe pump delivered the required reagent. After each addition, the system waited until thorough mixing and reaction completion. The pH value was recorded, and the algorithm generated a state vector to calculate the next reagent volume. When consecutive reagent additions involved different chemicals, the syringe pump was cleaned by flushing three times with 10 mL of water, with the waste solution discharged into a waste collection bottle, followed by a single rinse with 5 mL of the subsequent reagent to prevent cross-contamination. This procedure ensured precise pH control and minimized unwanted interactions between reagents.

Figure S1. Photographs of the automatic pH-adjustment device and its cleaning strategy during the dispensing of different reagents.

Section 2 Chemical Reagents

All aqueous stock solutions were made using Type I ultrapure water (18.2 MΩ•cm)

Base and Acid used for pH Adjustment

Hydrochloric acid, 36~38%, Sinopharm Chemical Reagent Co., HCl

Nitric acid 68%, Sinopharm Chemical Reagent Co., HNO3

Sodium hydroxide ≥99%, Sinopharm Chemical Reagent Co., NaOH

Buffer Solution for Calibration

Potassium acid phthalate, ≥99%, Shanghai Lei-Ci SA Scientific instrument Co., Ltd, KHC8H4O4

Mixed phosphate, ≥99%, Shanghai Lei-Ci SA Scientific instrument Co., Ltd, Na2HPO4 & KH2PO4

Sodium Tetraborate Decahydrate, ≥99%, Shanghai Lei-Ci SA Scientific instrument Co., Ltd, Na2B4O7⋅10H2O

Organic Solvent

Ethanol, ≥ 99.7%, Sinopharm Chemical Reagent Co., Ltd, C2H5OH

Diethyl Ether, ≥ 99.7%, Sinopharm Chemical Reagent Co., Ltd, (C₂H₅)₂O

Used for Mixed Acid

p-Hydroxybenzoic acid ≥99%, Sinopharm Chemical Reagent Co., C₇H₆O₃

Terephthalic acid ≥99%, Sinopharm Chemical Reagent Co., C₈H₆O₄ Acetic acid ≥99%, Sinopharm Chemical Reagent Co., CH₃COOH

5-Sulfosalicylic acid dihydrate ≥99%, Sinopharm Chemical Reagent Co., C₇H₆O₆S·2H₂O

Section 3 Experiment Validation

Experiment 1 Benchmarking Three Control Policies in a Robotic pH Adjustment Task

To evaluate the practical utility of three controllers in real buffer systems, we conducted a head‑to‑head comparison on an automated pH-adjustment platform under identical experimental conditions. At 25 ± 0.1 °C, the pH electrode was calibrated using a three-point calibration procedure with certified standard buffer solutions of pH 4.01, 6.86, and 9.18. Reagents for pH adjustment were prepared at the 250 mL scale: 0.1 M HCl by diluting ≈2.08 mL of ~12 M concentrated HCl to 250 mL with water; and 0.1 M NaOH by dissolving 1.00 g NaOH pellets in water and diluting to a final volume of 250 mL. Test buffers were also made at 250 mL: 0.05 M p‑hydroxybenzoic acid (pHBA; M = 138.12 g·mol⁻¹) by dissolving 1.73 g and making up to 250 mL (the clear supernatant was used during experiments due to limited solubility); 0.05 M terephthalic acid (TPA; M = 166.13 g·mol⁻¹) by dissolving 2.08 g to 250 mL (limited solubility—use the clear supernatant during experiments); 0.10 M acetic acid by adding 1.43 mL glacial acetic acid to water and diluting to 250 mL; and 14.7 mM 5‑sulfosalicylic acid dihydrate (SSA·2H₂O; M = 254.215 g·mol⁻¹) by dissolving 0.934 g and making up to 250 mL.

Mixtures were prepared according to the experimental labels:

Mixture 1: 0.05 M pHBA 6 mL + 0.05 M TPA 6 mL;

Mixture 2: 0.05 M pHBA 6 mL + 0.10 M acetic acid 6 mL;

Mixture 3: 0.05 M pHBA 6 mL + 14.7 mM SSA 6 mL;

Mixture 4: 0.05 M TPA 6 mL + 0.10 M acetic acid 6 mL.

Each mixture (12 mL) was combined in a three‑neck flask and mixed by gentle N₂ bubbling, then connected to the platform with the target set to pH 6.0 ± 0.1. Under strictly identical conditions, we ran three control modes in separate trials—rule‑based Bayesian algorithm, reinforcement learning (RL), and human judgement. The platform decided acid/base additions in real time and, after each dose, waited for the pH to stabilize before proceeding until the target criterion was met. We logged pH, dose type/volume, cumulative volume, step count, total duration, and overshoot for fair comparison of efficiency.

In parallel, to evaluate all of these pH adjustment algorithms, we conducted 3000 independent simulation experiments. All random processes were seeded with the same value (seed = 555) for reproducibility.

Each experiment was generated as follows:

1.1. The initial solution consisted of 11.0 mL of a 0.1 M polyprotic acid (monoprotic, diprotic, or triprotic).

2.2. The acid type was uniformly selected at random from {monoprotic, diprotic, triprotic}.

2.1.For monoprotic acids: a single pKa ~ U[2, 6].

2.2.For diprotic acids: pKa1 ~ U[2, 4], pKa2 ~ U[4, 7].

2.3. For triprotic acids: pKa1 ~ U[2, 4], pKa2 ~ U[4, 7], pKa3 ~ U[6, 8].

U[x1,x2] means uniform sampling from x1 to x2.

3.3. The initial pH was computed exactly by solving the charge balance equation for the pure acid solution at the given concentration and pKa values.

4. The target pH was uniformly sampled from [2, 11] and rounded to two decimal places.

5. To introduce realistic uncertainty and buffering effects unknown to the controller, three additional buffer species were present in the solution with:

5.1. pKa values independently drawn from U[2, 6],

5.2. The total molar amounts independently drawn from U[10-6, 0.5] mol.

Two sets of reagents were employed for pH adjustment: 0.1 M HCl/NaOH as the primary pair for coarse adjustment and 0.01 M HCl/NaOH as the secondary pair for fine adjustment. The system automatically switched to the secondary (dilute) reagents when pH oscillation occurs (three consecutive minimum-volume additions i.e. 0.01mL that crossed the target), thereby enabling finer control when needed. The maximum number of pH-adjustment steps was limited to 50, and success was defined as achieving a final pH within ±0.1 of the target. This procedure ensured broad coverage of adjustment curve shapes, buffering regimes, and adjustment directions (acidification or basification), providing a challenging and diverse benchmark for the algorithm.

Then we evaluated models in the simulation environment using three metrics: (1) experimental success rate, (2) average number of steps ± standard deviation, and (3) overshoot rate.

In the 3000 diverse simulated pH adjustment tasks, the Bayesian algorithm (success rate 94.23%, average steps ± standard deviation 12.81 ± 7.68, overshoot rate 42.70%), imitation learning (success rate 93.87%, average steps ± standard deviation 10.26 ± 7.21, overshoot rate 31.25%), and reinforcement learning (success rate 94.30%, average steps ± standard deviation 10.22 ± 7.24, overshoot rate 31.41%) all demonstrated excellent performance, far surpassing the traditional proportional-integral-derivative (PID) controller (success rate 29.67%, average steps ± standard deviation 30.40 ± 14.90, overshoot rate 24.38%).

Imitation learning and reinforcement learning led to efficiency, requiring only approximately 10.2 steps to complete the adjustment, while maintaining overshoot rates around 31%, achieving the best balance among success rate, efficiency, and safety. The Bayesian algorithm exhibited comparable success rates but slightly higher average steps (12.73) and a higher overshoot rate (42.70%). Although the PID controller had the lowest overshoot rate, its extremely low success rate and highly unstable step counts (standard deviation as high as 14.90) rendered it the poorest performer in this complex and variable titration environment.

Overall, imitation learning and reinforcement learning delivered the best comprehensive performance in this benchmark, followed by the Bayesian algorithm, while traditional PID control proved inadequate for handling polyprotic acid systems with unknown buffering interference.

Experiment 2 Wastewater Neutralization

In this experiment, we adopted a more flexible approach to pH adjustment through model guidance; the maximum reagent volume was capped at 5.0 mL, and any recommended value exceeding this limit was truncated. Before the experiment, the pH values of the acid and base used for adjustment were measured at 1.12 and 12.98, respectively. Consequently, a correction factor (approximately 1.26 in this experiment, indicating that theoretically 1.26 units of acid are required to neutralize 1 unit of base) was applied to the actual volume of the acid added to account for its impact on the final pH, based on the measured acid and base concentrations.

In this experiment, three samples of laboratory wastewater were collected, and their pH was adjusted to 7.0 ± 0.1. In this experiment, inorganic wastewater containing multiple transition-metal ions was collected on May 12, 2025, from a teaching lab in Xiamen University (Lujiaxi Building, Room 225). The pH electrode was calibrated at 25 °C using standard pH 4.01, 6.86, and 9.18 buffer solutions. Stock solutions of 0.1 M HCl and 0.1 M NaOH were connected to the dosing pumps of the automatic pH-adjustment device. For each run, 3 mL of the wastewater sample was transferred into a three-neck flask and diluted with 12 mL of ultrapure water (a total volume of 15 mL, corresponding to 5× dilution). High-purity N₂ was gently bubbled through a gas-dispersion frit to ensure mixing. The pH electrode was immersed so that the sensing bulb was fully covered by the solution, and data logging was initiated. The controller was set to maintain the pH at 7.0 ± 0.1. The reinforcement learning policy then autonomously dosed 0.1 M NaOH (base) or 0.1 M HCl (acid) as needed, with a stabilization time (20 s) allowed after each addition before the next dosing decision. Throughout the process, the system continuously recorded time, pH, reagent type/volume, and cumulative volume. The task was regarded to be complete once the solution reached and stabilized at pH 7.0 ± 0.1, at which point the total acid/base consumption and number of dosing steps were documented. The neutralized solution was subsequently collected and disposed of according to standard laboratory waste management procedures.

Experiment 3 Isolation of Casein by Adjusting pH to the Isoelectric Point

Yili Milk (date: 14 Feb 2025; batch B18 0103) was used in this experiment. From 12 mL of milk, the dry-mass yield of casein was typically 171 mg (14.25 mg·mL⁻¹). Figure S2 shows photographs of the separated casein.

Before pH-adjustment, the pH electrode was calibrated at 25 °C using standard pH 4.01, 6.86, and 9.18 buffer solutions. For each run, 12 mL of fresh milk was transferred into a three-neck flask and gently mixed either by low-flow N₂ bubbling through a gas-dispersion frit. The pH electrode was immersed so that the sensing bulb was fully covered, and the controller was set to maintain the pH at 4.6 ± 0.1. The reinforcement-learning routine was then initiated, allowing the system to autonomously dispense 0.1 M HCl (acid) or 0.1 M NaOH (base) as needed. After each dose, the system waited for stabilization before the next dosing decision, while continuously logging time, pH, reagent type/volume, and cumulative volume. Once the pH reached and stabilized within 4.6 ± 0.1, it was held at this value for 5 min to promote casein aggregation. The mixture was subsequently transferred to centrifuge tubes and centrifuged at 10,000 rpm for 12 min, after which the supernatant was discarded and the casein pellet collected. To remove residual fat, the pellet was washed with ethanol: diethyl ether (1:1, v/v), centrifuged again at 10,000 rpm for 12 min with the supernatant discarded; this defatting wash was repeated twice or until the supernatant appeared clear and non-opalescent. The pellet was then dried at 80 °C to constant weight, defined as two consecutive weighings at least 30 min apart with a mass difference < 1 mg. The dry mass was recorded to calculate the yield. As a reference, a typical run yielded 171 mg of casein from 12 mL of milk. Protein identity was confirmed by the Biuret test: dissolution of a small portion of the dried casein in 3% NaOH followed by addition of Biuret reagent produced a violet color, indicating the presence of peptide bonds. All ethanol/ether mixtures and acid/base wastes were disposed of in accordance with chemical-waste regulations.

Figure S2. Isolated casein obtained by pH adjustment to the isoelectric point and centrifugation.

Experiment 4 Determination of the Stability Constant for the SSA–Cu(II) Complex

In this experiment, we adopted a more flexible approach to pH adjustment through model guidance, where the maximum reagent volume was capped at 5.0 mL; any recommended value exceeding this limit was truncated. Before the experiment, the pH values of the acid and base used for adjustment were measured at 1.12 and 12.98, respectively. Consequently, a correction factor (approximately 1.26 in this experiment, indicating that theoretically, 1.26 units of acid are required to neutralize 1 unit of base) was applied to the actual volume of the acid added to account for its impact on the final pH, based on the measured acid and base concentrations.

For the determination of the SSA–Cu(II) complex stability constant, 0.888 g of Cu(NO₃)₂·3H₂O and 0.934 g of SSA·2H₂O were each dissolved in water and diluted to 250 mL, yielding two 14.7 mM stock solutions according to Table S1. These solutions were mixed in different volume ratios according to Table S2, transferred sequentially into a three-neck flask, and subjected to automated pH adjustment. The pH electrode was calibrated at 25 °C with standard pH 4.01, 6.86, and 9.18 buffer solutions. The device, loaded with 0.1 M HNO₃ and 0.05 M NaOH, was set to maintain pH 4.2 ± 0.1, with stabilization after each dosing step. Once the target pH was reached, the solution was kept quiescent or gently stirred for 5–10 min; if no noticeable pH drift was observed, UV–Vis measurements were performed. Absorbance spectra were collected in 1 cm cuvettes from 300 to 800 nm. The absorption at 400 nm (A₄₀₀), corresponding to the Cu–SSA complex band, was recorded, and a Job’s plot of A₄₀₀ versus SSA mole fraction (X_SSA) was constructed. A maximum near X_SSA ≈ 0.5 confirmed a 1:1 stoichiometry for the Cu-SSA complex. To determine the conditional stability constant, piece-wise linear regression of the Job’s plot provided the intersection absorbance A′, representing the maximum complex concentration. The degree of dissociation was then estimated as α = (A′ − A)/A, where α corresponds to the fraction of uncoordinated species. For a general equilibrium M + nR ⇌ MRₙ, the concentrations near the stoichiometric point can be expressed as [MRₙ] = c(1-α), [M] = cα, and [R] = ncα, with c as the total analytical concentration. The conditional stability constant is given by K′ = [MRₙ]/([M][R]ⁿ) = (1-α)/((nc)ⁿ αⁿ⁺¹). For the dominant 1:1 complex (n = 1), this reduces to K′ = (1-α)/(cα²). Using A′ and the derived α values, K′ was calculated to be approximately 6669 (mol/L)-1 at pH 4.2, indicating a strongly bound SSA–Cu(II) complex under mildly acidic conditions.

Table S1 Solutions Used in the Copper-Sulfosalicylic Acid Complex Formation

| reagent | mass | volume | concentration |
| --- | --- | --- | --- |
| Cu(NO3)2·3H2O | 887.9 mg | 250 mL | 14.7 mM |
| 5-sulfosalicylic acid·2H2O | 934.2 mg | 250 mL | 14.7 mM |
| NaOH | 500 mg | 250 mL | 0.1M |
| nitric acid | 1.58 mL | 250mL | 0.1M |

| Table S2 Solution Formulation of Experiment 4 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Solution ID | VCu (mL) | Ccu (mM) | VSSA (mL) | CSSA (mM) | (VSSA)/( VCu+ VSSA) | Vtotal (mL) |
| 1 | 10.10 | 9.90 | 0.92 | 0.90 | 0.08 | 11.02 |
| 2 | 9.18 | 9.00 | 1.84 | 1.80 | 0.17 | 11.02 |
| 3 | 8.27 | 8.10 | 2.76 | 2.70 | 0.25 | 11.02 |
| 4 | 7.35 | 7.20 | 3.67 | 3.60 | 0.33 | 11.02 |
| 5 | 6.43 | 6.30 | 4.59 | 4.50 | 0.42 | 11.02 |
| 6 | 5.51 | 5.4 | 5.51 | 5.4 | 0.50 | 11.02 |
| 7 | 4.59 | 4.50 | 6.43 | 6.30 | 0.58 | 11.02 |
| 8 | 3.67 | 3.60 | 7.35 | 7.20 | 0.67 | 11.02 |
| 9 | 2.76 | 2.70 | 8.27 | 8.10 | 0.75 | 11.02 |
| 10 | 1.84 | 1.80 | 9.18 | 9.00 | 0.83 | 11.02 |
| 11 | 0.92 | 0.90 | 10.10 | 9.90 | 0.92 | 11.02 |

Section 4 Bayesian Updating and Reinforcement Learning

Bayesian Parameter Updating

To infer the acid-base composition of the solution during pH-adjustment, we use an iterative Bayesian updating procedure based on parameter vector filtering. The unknown chemical system is represented as a set of buffer pairs (a set of monoprotic acids). For the i-th buffer pair, the parameters to be inferred are:

1. The acid dissociation constant of the i-th buffer pair pKa,i (dimensionless),

2. The total molar amount of the i-th buffer pair ntotal,i = nHA,i+ nA-,i (in mol).

The Bayesian updating procedure contains five steps: sampling, equilibrium prediction, likelihood calculation, resampling, and posterior updating.

1. Sampling

At each step, the current prior distributions are defined for every buffer pair. For buffer pair i, independent Gaussian priors are maintained:

pKa,i ~ N(upKₐ,ᵢ,σpKₐ,ᵢ)

ntotal,i ~ N(untotal,i,σntotal,i)

A set of M parameter vector, i.e., the complete parameter set to describe the system (indexed by j), is sampled 1000 times from these priors:

{θj}j=1M=(pKa,1j,ntotal,1j,…,pKa,ij,ntotal,ij)j=1M

Each parameter vector θj represents one complete candidate parameter set of the system, defining the corresponding pKₐ and total buffer amount for all buffer pairs. The ensemble of parameter vector collectively represents the current parameter distribution of the unknown solution chemistry.

2. Equilibrium prediction

Each parameter vector is then propagated through the equilibrium model to predict the pH after the current operation. For parameter vector θj, the prediction is based on chemical equilibrium and charge balance.

For the i-th buffer pair, an effective acidity constant pKa,ieff is first computed using Eq. S1. This quantity is obtained by combining the current inferred value pKa,i with a reference value pKa,iref.

pKa,ieff=pKa,iref+ wipKa,i- pKa,iref   where wi= wmax(1- tanh(k*σi))   (S1)

Where tanh is the hyperbolic tangent function, and the quantity pKa,iref is the reference acidity constant assigned to buffer pair i before online updating. pKa,iref is set by the user and serves as the chemically reasonable anchor for that buffer pair. In practice, pKa,iref is taken from the prior chemical model used to initialize the pH-adjustment problem, for example, from known or assumed acid-base chemistry of the solution, from literature values, or from the prior mean used at the beginning of the inference procedure. Its role is not to replace the inferred value, but to stabilize the prediction when the posterior for that buffer pair remains broad. wi∈0,wmax is the trust weight for buffer pair i. wmax is the maximum allowable weight set to a constant depending on the reliability of the prior information(here set to 0.2), and σi is the current posterior standard deviation of pKa,i. Because pKa is dimensionless, σi is also dimensionless. When σi is small, the inferred pKa,i is trusted more. When σi is large, the calculation is pulled closer to pKa,iref. This makes the equilibrium prediction more conservative when the inferred chemistry is still uncertain.

Using the resulting pKa,ieff together with the parameter vector composition, the hydrogen-ion concentration H+ is obtained by numerically solving the charge-balance equation in Eq. S2:

fpH=H++cNa+Na-KwH+-icA-,i-cCl-=cAk=1nk*j=1kKajH+k*D-cHCl=010-pH+cNa-Kw10-pH-cAi=1ni*Hn-iAi--cHCl=    (S2)

Here, H+ OH-is expressed in mol/L, with OH-=Kw/H+, and Kw=10-14 at room temperature. The predicted pH for the parameter vector θj is then calculated as pHpredj=-log10H+.

3. Likelihood calculation

After the experimental pH value pHobs is measured, each parameter vector is assigned a likelihood according to how well its predicted pH matches the observation. Assuming normally distributed measurement noise with a standard deviation σ=0.01 pH units, the likelihood is computed using Eq. S3:

L(θi)=NpHobs丨pHi, σ2=12πσexp-pHobs-pHi22σ2       (S3)

Parameter vector whose predicted pH values are closer to the measured pH receive larger likelihood weights.

4. Resampling

The likelihood weights L(θi) are normalized among the parameter vectors and used for multinomial resampling. Specifically, new parameter vectors are drawn with replacement from the original parameter vector set, and the sampling probability of each parameter vector is proportional to its normalized likelihood weight. In this way, parameter vectors that are more consistent with the measured pH are preferentially retained, whereas parameter vectors that are inconsistent with the measured pH gradually disappear from the ensemble.

The result of this step is an empirical approximation to the posterior distribution pθi∣pHobs. Importantly, the algorithm does not update parameters by moving one best-fit parameter vector directly. Instead, it updates the whole parameter vector population so that the parameter vector ensemble itself becomes a numerical representation of the posterior distribution.

5. Posterior updating

After resampling, the posterior statistics of each parameter are computed from the resampled parameter vector ensemble. For example, for buffer pair i, μpKaposti(the posterior mean of pKa,i) and σpKaposti (posterior standard deviation of pKa,i) are estimated using Eq. S4.

μpKaposti=1Mj=1MpKaij'; σpKaposti=1Mj=1M(pKa,ij'-μpKaposti)2  (S4)

The same calculation is performed for ntotal,i. These posterior means and posterior standard deviations are then used as the prior means and prior standard deviations for the next step. Thus, parameter updating proceeds recursively from one step to the next.

The Bayesian algorithm does not infer the chemical parameters only once at the beginning of the experiment. Rather, after every new pH measurement, the posterior distribution is updated, and the updated posterior becomes the starting prior for the next step. As the pH-adjustment proceeds, the posterior distributions usually become narrower and more informative, so the model gains increasing confidence in the chemical state of the system. This progressive refinement improves both pH prediction and subsequent volume recommendation.

Determining pH-Adjustment Volume Using Bayesian Parameter Updates and Control Logic

Once the posterior distributions have been updated at the current step, the controller converts the inferred chemical state into the next dosing command. In other words, the Bayesian module estimates the composition of the solution and its uncertainty, and the control module uses that information to decide how much the reagent should be added next.

The logic of this step is as follows. The controller first calculates the expected reagent volume that would ideally bring the solution to the target pH according to the current posterior model. It then moderates that ideal correction using quantities that describe how far the system is from the target, how strongly the solution responded to the previous addition, how uncertain the inferred chemistry remains, and how strongly buffered the current solution is. The final recommendation is therefore a chemically informed and uncertainty-aware dosing volume, rather than a purely empirical correction.

To match the practical resolution of the liquid-handling system, the candidate addition volume is discretized in steps of 0.01 mL over the range 0.01-9.99 mL. The final recommended volume is reported in mL, but whenever concentrations must be recalculated in the equilibrium model, the corresponding volume is converted to L.

The recommended reagent volume is determined from five quantities: 1. Current pH error magnitude (ΔpHerr); 2. Observed pH response to the previous addition (propagated with α through a term r); 3. Posterior uncertainty (propagated with α through a term u); 4. Effective buffering level (propagated with α through a term B); and 5. Theoretical required volume (Vreq).

The final recommended reagent volume Vadd is obtained by combining the five quantities above through the nonlinear mapping defined in Eq. S5:

Vadd=Vmin+Vmax-VmintanhαΔpHerr+0.1Vreq   (S5)

where Vmin=0.01 mL and Vmax=9.99 mL are the lower and upper operating bounds. The specific calculation methods for ΔpHerr, α(r,u,B) and Vreq will be discussed below.

1. Current pH error magnitude (ΔpHerr)

The first quantity is the absolute deviation of the current pH from the target pH:

ΔpHerr=pHcurrent-pHtarget.  (S6)

This quantity is dimensionless. It measures how far the current solution state is from the target. A larger ΔpHerr means that a stronger correction is required.

2. Observed pH response to the previous addition (r)

The second quantity describes how strongly the solution responded to the previous reagent addition. We define the previous-step pH response slope as

SpH=ΔpHprevVprev  (S7)

where ΔpHprev is the pH change caused by the previous pH-adjustment step and Vprev is the reagent volume added in that step, expressed in mL. Therefore, SpH has units of pH mL​-1.

A small value of SpH indicates that the solution responded weakly to the previous addition, so the next addition may be somewhat larger. A large value indicates that the solution is sensitive to reagent addition, so the next addition should be made more cautiously.

The response-related bonus term r is calculated from the previous-step pH response slope:

r=rbonus1-minSpHSpH,thresh,1   (S8)

where rbonus=0.5 is the maximum bonus and SpH,thresh=1.0 pH mL​-1 is the threshold above which no additional bonus is applied.

3. Posterior uncertainty (u)

The third quantity describes the current uncertainty in the inferred acid-base chemistry. We summarize this uncertainty by σpKa which is the average posterior standard deviation of the inferred pKa value. Because pKa is dimensionless, σpKa is also dimensionless. A larger value means that the inferred chemical model remains less certain, so the controller should behave more conservatively. The uncertainty factor is written as

u=minσpKaσmax,1,  (S9)

where σmax is a dimensionless normalization constant set to be 1.0.

4. Effective buffering level (B)

The fourth quantity is a bounded correction related to buffering. In the present framework, the mean of the posterior molar amounts across all buffer pairs is computed directly:

nbuffer=1Ni=1Nni,post    (S10)

where ni,post​ is the posterior mean molar amount of the i-th buffer pair (in mol), and N is the number of buffer pairs. This mean molar amount is then converted into the buffering factor B through Eq. S11:

B=clip1+0.1nbuffer-nref,min= 0.95, max=1.05   (S11)

where nref = 0.5 mol is the reference molar amount. The clipping in Eq. S11 constrains B to the range 0.95–1.05.

Here, B serves as a computationally efficient proxy for buffering strength. Rather than solving the full thermodynamic buffer capacity at each step, B captures the dominant physical trend: solutions with greater buffer content can typically absorb a larger reagent addition while producing only a modest pH shift. The clipping to the range 0.95–1.05 ensures that this correction remains conservative and does not override the equilibrium-based dosing logic.

5. Theoretical required volume (Vreq)

The fifth quantity is the theoretically required volume, denoted Vreq. This is the estimated reagent volume that would bring the solution exactly to the target pH according to the current posterior model. It is obtained by solving the nonlinear equation in Eq. S12:

Vreq: FVreq=0, FV=pHV-pHtarget  (S12)

Where pHV is the pH after adding liquid of volume V. The meaning of Eq. S12 is direct. For each trial volume V, the equilibrium model predicts the pH that would result after adding that amount of reagent. The correct value Vreq is the one for which the predicted post-addition pH is equal to the target pH. In this sense, the controller solves a one-dimensional inverse equilibrium problem.

The function pHV is evaluated in two steps.

Step 1: effective acidity constants.

For each buffer pair, an effective acidity constant is calculated from the posterior mean and the reference value using the uncertainty-weighted rule in Eq. S1.

Step 2: post-addition equilibrium prediction.

After a candidate reagent volume V is specified, the amount of strong acid or strong base introduced by the reagent is first calculated from the concentration and the added volume. The molar amounts of the chemical species in solution are then updated to account for the addition. At the same time, the total solution volume is increased to the post-addition volume. Because the concentrations change both through neutralization and through dilution, all updated molar amounts are converted to post-addition concentrations using the new total volume. The equilibrium calculation is then solved again using the updated composition and the effective acidity constants from Step 1. The resulting hydrogen-ion concentration gives the predicted post-addition pH, denoted pHV.

Thus, Vreq is the model-based one-step correction implied by the current posterior chemical model.

Recall the Eq. S5, where the volume added is:

Vadd=Vmin+Vmax-VmintanhαΔpHerr+0.1Vreq   (S5)

In Eq. S5, the scaling factor α is defined by Eq. S13:

α=Φ1+r1-0.1uB  (S13)

The role of α is to control how aggressively the controller follows against Vreq. When the chemical model is more reliable and the solution response is mild, the controller can move more directly toward the target with larger α values. When the posterior remains uncertain, or the solution responds sharply to reagent addition, the controller behaves more conservatively.

The baseline coefficient Φ is a dimensionless constant that sets the overall response scale of the controller, and we set Φ=0.2.

Taken together, these terms ensure that the final recommended volume is guided mainly by the equilibrium-based correction Vreq, but is reduced when the current chemistry is uncertain or when the solution has shown strong pH sensitivity.

Imitation Learning

To provide a strong initialization for reinforcement learning, we first trained a supervised imitation-learning policy using successful trajectories generated by the Bayesian controller. A total of 8,000 successful pH-adjustment trajectories were simulated. Each trajectory consists of a sequence of state-action pair st,at, where t denotes the index of pH-adjustment step.

At step t, the state vector is defined as

st∈R5,

and contains the following five features:

- the current measured pH, pHt (dimensionless);

- the target pH, pHtarget (dimensionless);

- the change in measured pH between the current and previous steps,

ΔpHt=pHt-pHt-1,   (14)

- which is also dimensionless;

- the current pH difference toward the target,

et=pHtarget-pHt,   (15)

- which is dimensionless; and

- the reagent volume used in the previous step, Vt-1, expressed in mL.

Thus, the state representation combines the current system status, the control objective, the recent pH response, and the most recent dosing magnitude.

The action at corresponds to the decision taken at step t, including both the reagent type (acid or base) and the dispensing volume. In the implementation of imitation learning, the volume component is discretized into 1000 categories from 0.01 to 10.00 mL with a resolution of 0.01 mL. Therefore, the policy is trained as a classification model over a finite action set of candidate dosing volumes.

All successful trajectories were decomposed into individual state-action pairs, randomly shuffled, and split into training (70%), validation (15%), and test (15%) subsets. This preprocessing removes temporal ordering during supervised fitting and allows the model to learn the mapping from local pH-adjustment state to dosing action.

The imitation-learning policy is a multilayer perceptron (MLP) with a 5→256→256→1000 architecture. The input layer receives the 5-dimensional state vector st. Two hidden layers with 256 neurons each and Rectified Linear Unit (ReLU) activation are used to model nonlinear relationships between pH state variables and dosing decisions. The output layer produces a 1000-dimensional logit vector, whose components correspond to the discrete candidate volumes from 0.01 to 10.00 mL.

During inference, the predicted reagent volume is selected as the category with the highest logit (equivalently, the highest predicted probability after softmax), and then mapped back to its corresponding physical dosing volume in mL. In this way, the network outputs a physically interpretable control action rather than an abstract latent quantity.

The model was trained using cross-entropy loss and the Adam optimizer with a learning rate 1×10-3 for 80 epochs. On the held-out test set, the trained model achieved an MSE of 0.0062, an MAE of 0.0335, and an R2 value of 0.9992, indicating that the network learned to reproduce the Bayesian controller with high fidelity. This imitation policy was then used as the initialization for subsequent reinforcement learning.

Optimizing Neural Network Parameters via Reinforcement Learning

Building on the imitation-learning initialization, we further optimized the policy using reinforcement learning. Specifically, we employed the REINFORCE algorithm, a Monte Carlo policy-gradient method, to improve decision-making through direct interaction with the environment. The objective is to maximize the expected cumulative reward by adjusting the policy parameters θ.

In the reinforcement-learning stage, one episode corresponds to a complete pH-adjustment experiment, starting from an initial solution state and ending when either the target pH is successfully reached or the maximum allowed number of steps is exceeded. At each step t, the policy πθat∣st receives the current state st and outputs an action at, i.e., the reagent type and dispensing volume.

The trajectory collected in one episode is written as

τ=s1,a1,r1,s2,a2,r2,…,sT,aT,rT,

where rt is the reward obtained after taking action at in state st, and T is the terminal step of the episode.

Parameter update process

The REINFORCE 1 update consists of four steps.

1. Trajectory sampling

A complete trajectory τ is sampled by executing the current policy in the simulated pH-adjustment environment until successful completion or termination at the maximum step limit.

2. Discounted return calculation

For each step t, the discounted cumulative return Gt is computed according to Eq. S16:

Gt=rt+γGt+1+γ2Gt+2+…      (S16)

Here, rt denotes the reward received immediately after taking the action at step t. The return Gt includes not only the immediate reward rt, but also all future rewards obtained at subsequent steps t+1,t+2,…,T until the end of the trajectory, discounted geometrically by the factor γ. In this work, γ=0.99, so rewards received later in the episode still contribute strongly, but slightly less than immediate rewards. Since reward values are dimensionless by construction, Gt is also dimensionless. The large value of γ emphasizes successful completion of the full pH-adjustment trajectory while still weakly favoring earlier achievement of the target.

3. Reward normalization/clipping

Before computing the return, the instantaneous rewards are clipped to a predefined range ([-4.1, 4.1], see below for more details) to prevent unstable updates caused by excessively large positive or negative values. This step improves numerical stability during policy-gradient optimization.

4. Policy-gradient update

The network parameters θ are then updated using the REINFORCE gradient estimator to maximize the expected return Jθ. In practice, after each sampled episode, the cumulative policy loss is backpropagated, and the model parameters are updated using the Adam optimizer with a learning rate 1×10-3.

After multiple rounds of online interaction and parameter updates, the policy is no longer constrained to imitate the Bayesian controller exactly. Instead, it can learn action patterns that improve task success and dosing efficiency under the reward design of the environment.

Reward Function Design

The reward at step t, denoted rt, is constructed from six components:

rt=clip(rtdense+rtstep+rtovershoot+rtreagent+rtvolume)+rtterminal. (17)

All reward terms are dimensionless. The purpose of this composite reward is to balance three control objectives simultaneously: (1) reaching the target pH accurately (rtdense, rtterminal), (2) doing so in as few steps as possible (rtstep), and (3) avoiding unstable or chemically unreasonable dosing behavior (rtovershoot, rtreagent, rtvolume). The clip part (rtdense+rtstep+rtovershoot+rtreagent+rtvolume) is clipped to [-4.1, 4.1]

1. Dense reward(rtdense)

The dense reward (rtdense) encourages the policy to reduce the pH error at every step. It is defined in Eq. S18:

rtdense=λet-1-et1+tM-ttM  (S18)

In Eq. S11, et denotes the pH error at step t, defined as

et=pHtarget-pHt, (S19)

which is dimensionless. The step index t ranges from 1 to tM, where tM is the maximum allowed number of steps and is fixed at 50 in this work. Accordingly, tM-t denotes the number of remaining steps after the current step t, and the ratio tM-ttM provides a normalized weighting factor that assigns greater importance to reducing the pH error at earlier stages of the adjustment process. The scaling factor λ is set to -0.03. This term provides dense feedback throughout the episode and rewards gradual movement toward the target pH rather than only final success.

2. Step penalty

A fixed penalty(rtstep) is applied at every step to discourage unnecessarily long pH-adjustment sequences. This term encourages the controller to complete the task efficiently and reduces dithering near the target. After optimizing hyperparameters via grid search, it was found that setting this parameter to 0 yields the best results.

3. Overshoot penalty

An overshoot penalty(rtovershoot) is applied whenever the pH deviation exceeds the tolerance threshold of 0.1 pH units. The corresponding functional form is given in Eq. S20:

rtovershoot=-w*σσet-δ,if et-1•*et<0 and et<>0.1 0,   otherwise   (S20)(S20)

Here, σx=(1+e-x)-1 is the sigmoid function et is the pH deviation from the target, w is a fixed weighting coefficient set to 0.2, and δ=0.1 is the overshoot boundary. Since pH is dimensionless, both et and δ are dimensionless. This term penalizes aggressive actions that drive the system too far beyond the target and therefore promotes smoother, safer adjustment behavior.

4. Incorrect reagent penalty

A fixed penalty(rtreagent) is assigned when the selected reagent direction is chemically inconsistent with the control objective. Specifically, this occurs when acid is added while the controller should increase pH, or when base is added while it should decrease pH (defined as "wrong reagent used"). This term encodes a basic chemical prior into the reward and helps prevent obviously wrong dosing decisions. Here α is set to 1, and et is the pH deviation from the target.

rtreagent-αet,    wrong reagent used0,   otherwise(S21)

5. Overshoot-related volume penalty and bonus

The reward also includes a volume-sensitive term associated with overshoot behavior (rtvolume). If the previous action caused overshoot, a penalty is applied based on the added volume. The detailed formulation is given in Eq. S22:

rtvolume=-pvol×vt+bvol*maxVover-vt,0  (S22)

In this expression, vt is the volume dispensed at step t, expressed in mL; pvol is a fixed penalty coefficient set to 0.1; bvol is a fixed bonus coefficient also set to 0.1; and Vover is the volume of the action that caused the overshoot, max refers to the maximum function. If overshoot occurred in the previous step and the next action uses a smaller volume than the overshooting action, the controller receives a positive bonus. This design encourages adaptive correction: once the policy detects that it has been too aggressive, it is rewarded for becoming more conservative in the following step.

6. Terminal reward

A terminal reward(rtterminal) is assigned when the episode ends successfully. The functional form is given in Eq. S23:

rtterminal=2B,if  t<tM2B,if  t≥tM20,otherwise (S23)

A larger terminal reward is granted when the task is completed in fewer than half of the maximum allowed steps. In Eq. S23, B is a fixed reward constant set to 3.9. This term explicitly values both success and efficiency, thereby encouraging the policy not only to reach the target pH but to do so promptly.

Taken together, the reward function combines local shaping terms and final task-level incentives. The dense reward and overshoot terms guide short-term behavior, the step penalty promotes efficiency, the incorrect-reagent penalty imposes chemically meaningful directional constraints, and the terminal reward emphasizes overall task completion.

Training Procedure

For reinforcement learning, the policy network initialized by imitation learning was further optimized over 500 training episodes. In each episode, a random pH-adjustment problem was generated from a family of monoprotic, diprotic, and triprotic acid systems. The acid concentration was fixed at 0.1 M, and target pH values were randomly sampled from 2 to 11. The maximum number of steps was 50.

After training, the updated policy was evaluated on 3000 test experiments generated using the same random seed as the Bayesian-parameter-updating benchmark. In a more challenging setting using only 0.1 M HCl and 0.1 M NaOH, the reinforcement-learning policy achieved a success rate of 87.27% with an average of 7.88 steps among successful experiments, whereas the Bayesian baseline achieved a success rate of 77.63% and required an average of 10.20 steps. Thus, the learned policy improved the success rate by 9.64 percentage points and reduced the average number of steps by 22.7% relative to the Bayesian controller.

In the broader 3000-task simulation benchmark, the reinforcement-learning controller achieved a success rate of 94.30%, an average step count of 10.221±7.241, and an overshoot rate of 31.41%, performing comparably to imitation learning and better than the Bayesian controller in efficiency. These results indicate that reinforcement learning preserves the strong baseline behavior learned from imitation while further improving control efficiency through direct optimization of task-level reward.

The raw code for the machine learning process is available at GitHub https://github.com/Wang-Group/pH-Titration.

Reference

1.	R. J. Williams, Machine Learning, 1992, 8, 229–256.